2θ − ω scan using PDS and PIXcel in Bragg-Brentano geometry

(X’Pert_2 only)

I. Login

- Start Data Collector.

- Type your UserName and Password.

- Select Instrument – Connect.

- Choose Configuration PDS + PIXcel.

- Click OK.

II. Hardware Setup

- X-ray Tube is in Line Focus.

- Goniometer Resolution set to “Normal 0.001 deg”

- Incident Beam Optics – PDS (Programmable Divergence Slit).

Note: if you have to change incident beam optics please first turn Automatic attenuator to “Activate” status and then unplug attenuator cable.

  - Insert Ni filter.

  - If sample is smaller than 25 mm, insert correct size Mask.

- Diffracted Beam Optics – PIXcel.

- Insert Soller Slit.

III. Data Collector Software

- Select the Incident Beam optics tab.

  - Double click any item. Incident beam optics window will appear.

  - Go through all tabs and select proper optic components.

  - Divergence Slit – “Fixed” mode and slit set to 1/32o.

  - Set Automatic Attenuator Usage to “Do not switch” and “Activated”.

- Select the Diffracted Beam optics tab.

  - Double click any item. Diffracted beam optics window will appear.

  - Go through all tabs and select proper optic components.

  - Detector – “Receiving Slit” mode and active length set to 0.16 mm.

- Select Instrument Settings tab.

  - Double click any item in the tree view to prompt another window.

  - Press X-ray tab. Set generator power to 45 kV and 40 mA.

IV. Diffractometer Zero Alignment

- In Instruments Settings check Z position. If it is larger than 5 mm move it back to at least 5 mm.

- Move all other motors to zero positions.

- From Menu Select Measure – Manual Scan.

- From the Scan Axis drop down menu select 2Theta.

- Enter Range 1o, Step Size 0.005o, and Time per Step 0.1sec. Then press Start.

- After scan is finished, right click on mouse and select Move Mode. Move Scan Axis to the center of the mass of the peak.

- In User Settings – Sample Offsets set current 2Theta position to zero.

- Note the direct beam intensity.

V. Sample Mounting

- Mount sample using scotch tape. If the sample is large, supplied clips can be used to hold the sample.

- If in the Instrument Settings tab X = 0.0 and Y = 0.0, beam is positioned at the center of a sample stage (aluminum disk).

VI. Moving Sample into the Beam Position

- Using supplied Micrometer.

- Mount micrometer onto the MRD cradle. Close the doors.

- In the Instruments Settings window, double click any item in the tree view to prompt another window.

- Move Z until micrometer reads (1.00±0.02) mm. This is the correct sample height.

- Using direct beam.

- Note the direct beam intensity. In Instruments Settings move Z to higher values until intensity starts to drop.

- In Manual Scan window from the Scan Axis drop down menu select Z. Enter Range 2 mm, Step Size 0.01 mm, and Time per Step 0.2 sec. Then press Start.

- After scan is finished, right click on mouse and select Move Mode. Move Scan Axis to the intensity value corresponding to ½ of the direct beam intensity.

- Z is aligned. Close shutter.

Note: sample height alignment using direct beam will work correctly if sample size is larger than the beam size in axial direction.

VII. Measurement

- Select Incident Beam optics tab.

- The Divergence Slit can be used in “Fixed” and “Automatic” modes. If “Fixed” mode is set, select the divergence of the slit to desired value from 1/32o to 4o. If “Automatic” mode is set, select a proper irradiated length of the sample.

- Set Automatic Attenuator Usage to “Do not switch” and “Deactivated”.

- Click on Diffracted Beam optics tab.

- Select Detector tab. Set detector mode to “Scanning” and active detector angle set to 2.511o.

- Simplest way to execute scan is to do a Manual Scan. It is a relative scan i.e. executed around current goniometer position with the range specified in Manual Scan window.

- To do 2Theta – Omega scan first move 2Theta and Omega to middle positions of the scan range.

- In Manual Scan window select Scan Axis “2Theta – Omega” and appropriate Range, Step Size and Time per Step.

- Click Start. When scan is completed, save it through File – Save As menu. Manual Scan will be lost if it is not saved.

Note: In symmetrical scan mode Omega axis is always ½ of the 2Theta axis.

VIII. Logging out

- Close the shutter.

- Switch Automatic Attenuator to “Do Not Switch” and “Activated”.

- Move all angles to zero positions and Z to 5 mm.

- Lower the power of the x-ray tube to 40 kV and 10 mA.

- Close Data Collector.