## Slide 1: Frequency Doubling with Nd:YAG laser
Tony Hyun Kim, Ilan F. Almog
12/9/2008
6.161 Modern Optics Lab Project

## Slide 2: Topics to be discussed
Nd:YAG laser restoration
  Principles of operation
  Repair of the cooling system, replacement of filter
  Fixing the deionizer cartridge
  Replacement of pumping lamp
Frequency-doubling experiment
  Brief theory
  Experimental setup; sources of error
  Results and conclusion

## Slide 3: Nd:YAG – Principles of operation
Nd:YAG crystal has infrared (1064nm) line.
Inversion achieved by Kr flash lamp.
System (lamp) must be water cooled.

## Slide 4: Our Nd:YAG, the Quantronix 116
Insights from previous repair attempts…
Plumbing issues:
  Various leaks
  Water filter too old
Laser head:
  Lamp malfunction

## Slide 5: 6.162: Modern Plumbing Laboratory

## Slide 6: 6.162: Modern Plumbing Laboratory
Two system-side (water) circuits:
  Left: Cools the laser
  Right: Self-cleans the water supply
Water filter was easily replaced, but…

## Slide 7: Deionizer cartridge fiasco
From
Pump
To
Reservoir

## Slide 8: Deionizer cartridge fiasco

## Slide 9: Deionizer cartridge: Solution
Tubes, clamps and epoxy  No leaks; perfect operation

## Slide 10: Replacement of the pump lamp
Suspicion of malfunctioning lamp confirmed
With plumbing and lamp fixed, laser operates!

## Slide 12: Conclusions (Nd:YAG repair)
No known leaks.
Have established protocols for maintenance:
  Refilling the reservoir.
  Wetting the pump for initial run.
  Pump electrical reset, and…
Observed up to 14W of IR power

## Slide 14: Freq. doubling introduction
E2 term in polarizability (e.g. KDP crystal)
Double 1064nm (IR)  Obtain 532nm (green)
Our experiment: Characterize the conversion efficiency as a function of input power.
Nonlinear crystal
Incident IR

## Slide 15: Experimental setup
Use dispersion of glass to achieve separation of IR (shown red) and green beams;
Can measure power of green beam only!

## Slide 16: Experimental setup
Use pick-off mirror to measure input IR power without disturbing other optics

## Slide 17: Experimental setup

## Slide 18
Generation of green light by freq. doubling with Quantronix 116

## Slide 19: Conclusions (Freq. doubling)
Observed frequency doubling using the newly-repaired Quantronix Nd:YAG laser.
Based on our alignment, obtained ~5.2% conversion efficiency.
  Efficiency independent of input IR power

## Slide 20: Concluding remarks
Showed that Nd:YAG system and KDP crystal are functional. Opens possibilities for future experiments.

## Slide 21: Acknowledgments
Mike the plumber;
Dave the optics guy;
The folks who shared the optics lab with us
  That is: Americo, Amrita, Matt, (Raymond);
And, of course, Bill (and Prof. Warde)